## Slide 1: Cross-calibration of SOHO/ERNE proton and heavy ion measurements with other particle instruments at L1
E. Valtonen, M. Paassilta, R. Vainio, O. Raukunen, and E. Riihonen
Space Research Laboratory
Department of Physics and Astronomy
University of Turku

## Slide 2: Contents summary
L1 instrument cross-calibrations within SEPServer project
ERNE instrument
Instrumental effects
Finding solutions
Conclusions

## Slide 3: L1 instrument cross-calibrations within SEPServer project
SEPServer
  EU FP7 funded project 2010-2013
  Server to access SEP and related EM data and analysis tools
Data quality assessment
  Cross-calibration of instruments at L1
  E.g., SOHO/ERNE, SOHO/EPHIN, ACE/EPAM, and ACE/SIS proton and helium spectra and SOHO/ERNE and ACE/SIS heavy ion spectra
  See also Poster Id 72832 by B. Heber et al.
  Generally relatively good agreement, but also some instrumental effects found

## Slide 4: “Relatively good agreement”

## Slide 5: ERNE instrument
Two sensors
Low Energy Detector (LED)
High Energy Detector (HED)
70 mm
Proton and helium spectra constructed onboard
Heavy ions as pulse heights to ground
D1
D2
D3
Four layers of strip detectors

## Slide 6: Instrumental effects
?
Steep fall?
?
Bump
Inaccurate G?
In addition:
Saturation of HED during largest SEP events
  Not considered here

## Slide 7: Finding solutions
Bump in proton spectrum
  Due to background: inefficient anticoincidence shield
Particles “stopping” in D3
Protons
Helium
Background

## Slide 8
Ground-processed
Selected ground-processed converted to onboard-analysed

## Slide 9: Finding solutions
Steep fall of proton spectrum
IMP-8 GME data from
http://dev.sepem.oma.be/index.php
  Partial cause: onboard event selection?
    Cuts out highest energy protons
  Instrument aging?

## Slide 10
GOES 11 SEM data from
ftp://satdat.ngdc.noaa.gov/sem/goes/
data/new_avg/2004/11/goes11/csv/

## Slide 11: Finding solutions
Geometric factor for heavy ions
  New calculation for ERNE/LED
  Geant 4 simulation fro ERNE/HED 16O
Updated G

## Slide 12
  For heavy ion measurements, see also poster in Session 8 by Raukunen et al.
Updated G

## Slide 13: Conclusions
Cross-calibration work of L1 particle instruments continued from the SEPServer project
SOHO ERNE proton spectrum investigated with the goal of improving match with other instruments
  Small SEP event/quiet time “bump” removal
  Steep spectral fall at energies 70-80 MeV
  Further work required
Heavy ion spectra with slightly revised geometric factors
  Geant 4 verification of geometric factors for all (heavy) elements